OSPAR Reference Method of Analysis for the Determination of the Dispersed Oil Content in Produced Water

Agreement 2005-15

Contents 	Page

1	Scope	3

2	Normative references 	3

3	Term and definition	3

4	Interferences	3

5	Principle	3

6	Reagents	4

7	Apparatus 	5

8	Sampling and sample preservation 	6

9	Procedure	6

10	Test report	9

11	Precision	10

Annex A (informative) Example of a column	11

Annex B (informative) Examples of gas chromatograms of mineral oil standard and water samples 	13

Annex C (informative) Determination of boiling range of a mineral oil from the gas chromatogram 	21

Bibliography 	21

# Determination of the content of dispersed mineral oil in produced water using gas chromatography with flame-ionization detection

# 1 	Scope

This document specifies the OSPAR reference method for the determination of dispersed oil in water by means of gas chromatography. The method is designed for produced water and other types of waste water discharged from gas, condensate and oil platforms and allows the determination of the dispersed oil content in concentrations above 0,1 mg/l. The method is a modification of ISO 9377-2 in order to include the determination of certain hydrocarbons with boiling points between 98 and 174°C.

NOTE 1 The mass concentration of animal and vegetable fat in the test sample should not exceed 150 mg/l, because at higher values the adsorption capacity of the clean-up column packing may not be sufficient.

NOTE 2 In the case of highly polluted waste water, especially if containing a high amount of surfactants, a loss in recovery may occur.

# 2 	Normative references

The following normative documents contain provisions which, through reference in this text, constitute provisions of this document. For dated references, subsequent amendments to, or revisions of, any of these publications do not apply. However, parties to agreements based on this document are encouraged to investigate the possibility of applying the most recent editions of the normative documents indicated below. For undated references, the latest edition of the normative document referred to applies. Members of ISO and IEC maintain registers of currently valid International Standards.

ISO 5667-3:1994, Water quality - Sampling - Part 3: Guidance on the preservation and handling of samples.

ISO 8466-1:1990, Water quality - Calibration and evaluation of analytical methods and estimation of performance characteristics - Part 1: Statistical evaluation of the linear calibration function.

# 3 	Term and definition

For the purposes of this document, the following term and definition applies.

  - Dispersed oil content by GC-FID

The sum of the concentrations of compounds extractable with n-pentane, not adsorbed on Florisil and which may be chromatographed with retention times between those of n-heptane (C7H16) and n-tetracontane (C40H82) excluding the concentrations of the aromatic hydrocarbons toluene, ethyl benzene and the three isomers of xylene.

NOTE Substances complying with this definition are long-chain or branched aliphatic, alicyclic, and (alkyl-substituted) polycyclic aromatic hydrocarbons, with boiling points between 98 and 525°C.

# 4 	Interferences

Compounds of low polarity (e.g. halogenated hydrocarbons) and high concentrations of polar substances may interfere with the determination.

Surface-active substances may interfere with the extraction step.

# 5 	Principle

The water is extracted with n-pentane. Polar substances are removed by clean-up on Florisil. The purified aliquot is analysed by capillary gas chromatography using a non-polar column and a flame ionisation detector (FID). The total peak area between n-heptane and n-tetracontane is measured. The peak areas of the aromatic hydrocarbons toluene, ethyl benzene and the three isomers of xylene are subtracted from the total peak area. The concentration of dispersed oil is quantified against an external standard consisting of a mixture of two specified mineral oils, and the content of dispersed oil is calculated.

It is absolutely essential that suitably qualified staff carry out the test described in this document.

It should be investigated whether and to what extent particular problems will require the specification of additional marginal conditions.

# 6 	Reagents

All reagents shall be reagent grade and suitable for their specific purpose. The suitability of the reagents and solutions shall be checked by carrying out a blank test.

6.1	Water for the preparation of solutions

Distilled water, or water from a generator of purified water capable of removing organic traces, for example using activated carbon, shall be used.

  - n-Pentane, C5H12

6.3	Sodium sulfate, anhydrous, Na2SO4.

6.4	Magnesium sulfate heptahydrate, MgSO4.7H2O.

6.5	Mineral acid, e.g. hydrochloric acid, c(HCl) = 12 mol/l (ρ = 1,19 g/ml).

6.6	Acetone, C3H6O

6.7 	Florisil, grain size 150 μm to 250 μm (60 mesh to 100 mesh), heated to 140 °C for 16 h and stored in a desiccator.

6.8	Mixture of mineral oils

6.8.1	Standard mixture

Weigh accurately equal amounts of two different types (type A and type B, both containing no additives) of mineral oil and add enough extraction solvent (6.11.2) to give a total hydrocarbon concentration of about 10 mg/ml.

Type A should show discrete peaks in the gas chromatogram. An example is diesel fuel without any additives. See EN 590 for further information. Type B should have a boiling range higher than that of type A and should have unresolved signals in the gas chromatogram. An example of this type is a lubricant without any additives, boiling range 325°C to 460°C.

6.8.2	Calibration mixture

Prepare at least five different calibration solutions by diluting aliquots of standard mixture (6.8.1) with the extraction solvent (6.11.2). The following concentrations may be suitable:

0 (blank), 0,2 mg/ml, 0,4 mg/ml, 0,6 mg/ml, 0,8 mg/ml and 1,0 mg/ml.

Higher concentrations may be advisable for other applications.

Store the calibration mixture tightly sealed in a refrigerator (4 °C to 8 °C). The calibration mixtures are stable for up to six months.

6.8.3	Quality Control (QC) standard

Prepare a standard solution according to 6.8.1 in acetone (see 6.6) with a mass concentration of e.g. 1 mg/ml. The exact concentration should be about a thousand times the desired application range.

NOTE	 When using a lubricant for QC, the stock solution is easily prepared in extracting agent (6.2), which is further diluted in acetone about tenfold before spiking the QC water.

Store the solution tightly sealed in a refrigerator (4 °C to 8 °C). It is stable for up to six months.

6.9 	Standard mixture of n-alkanes for system performance test

Dissolve n-alkanes with even carbon numbers (C20, C40 and at least three further n-alkanes) in extracting agent (6.2) to give concentrations of approximately 50 μg/ml of the individual components. It may be necessary to use a different solvent, e.g. heptane, for the first solution; in this case dilute this first solution with extracting agent (6.2).

Store the standard mixture tightly sealed in a refrigerator. It is stable for up to six months.

NOTE 1: This solution is used to verify the suitability of the gas chromatographic system for the resolution of n-alkanes as well as for the detector response.

NOTE 2: This solution is used to give information on the retention times of the n-alkanes to characterize the hydrocarbons in the sample.

6.10 	Reference compounds

6.10.1	n-Heptane, C7H16

6.10.2 	n-Tetracontane, C40H82

6.10.3	n-Eicosane, C20H42

6.10.4	Toluene, C7H8

    - Ethyl benzene, C8H10

    - o-Xylene, C8H10

    - m-Xylene, C8H10

    - p-Xylene, C8H10

6.11 	Extraction solvent containing reference compounds

6.11.1 	Extraction solvent stock solution.

Dissolve 20 mg of n-tetracontane (6.10.2) in n-pentane (6.2). Then add 20 (l of n-heptane (6.10.1) and dilute with n-pentane to 1000 ml.

Store the solution tightly sealed in a refrigerator. It is stable for up to six months.

NOTE 	n-Tetracontane is only moderately soluble in n-pentane. Slight warming or treatment with ultrasonics accelerates the dissolution process.

6.11.2		Extraction solvent standard solution.

Immediately prior to use, dilute the stock solution (see 6.11.1) with tenfold extracting agent.

6.12		Stearyl stearate test solution, C36H72O2

Dissolve 200 mg of stearyl stearate in 100 ml of extraction solvent standard solution (6.11.2).

NOTE 	This solution is used to check the efficiency of the clean-up procedure.

Store the solution tightly sealed in a refrigerator. It is stable for up to six months.

# 7 	Apparatus

7.1	Usual laboratory glassware

Clean all glassware by the usual procedures for this type of analysis and check for purity (measurement of blank). If necessary, rinse the glassware with extracting agent (6.2) and re-check the blank.

  - Gas chromatograph, equipped with a non-discriminating injection system and a flame ionization detector. The injection system shall allow injections up to a volume of at least 100(l.

7.3 	Column for gas chromatography, fused silica, with one of the following stationary phases:

non-polar, immobilized 100 % dimethylpolysiloxane, or 95 % dimethyl-/5 % diphenylpolysiloxane, or modified siloxane polymer and typical dimensions:

length:		5 m to 30 m

internal diameter:	0,25 mm to 0,53 mm

film thickness:	0,25 μm to 1,2 μm

For an example, see annex B.

It is recommended to use a precolumn (e.g. 2 m, 0,53 mm i.d., deactivated fused silica).

7.4	Data system, suitable for integrating the total area of the gas chromatogram and suitable for compensating for 'column bleeding' and for re-integrating after drawing a new baseline.

7.5	Sampling bottles, glass, with ground glass stopper, capacity 250 ml and 1 000 ml, or with screw cap, lined with PTFE (polytetrafluoroethene).

The sampling bottle shall allow direct extraction from the bottle.

7.6	Centrifuge.

7.7	Centrifuge tubes, of capacity 100 ml, with suitable (screw)cap.

7.8	Microseparator or other suitable device for phase separation.

7.9 	Clean-up columns, made from glass, with frit of sinter porosity 2, for example see Figure A.1.

7.10 	Magnetic stirrer with bar, of length suitable to ensure thorough mixing.

# 8	Sampling and sample preservation

Sampling and storage shall be in accordance with ISO 5667-3 and other OSPAR documents.

Fill the sampling bottle (7.5) to approximately 90 % with the sample, seal tightly and weigh (m1). Keep the sample at about 4 °C and extract the sample as soon as possible, but in any case within four days.

If necessary, preserve the sample by acidifying with mineral acid (6.5) to pH 2.

NOTE  If formation of emulsions or a concentration of animal and vegetable oils > 150 mg/l is expected; it is advisable to withdraw additionally a smaller sample volume in a 250 ml sampling bottle.

# 9 	Procedure

9.1 	Blank test

Carry out blank tests with each series of tests in accordance with 9.3, including all reagents and glassware in the same way as the samples.

9.2 	Determination of the recovery

Determine the recovery at regular intervals, preferably in each series of tests, using 900 ml of water (6.1) to which 1,0 ml of the QC standard (6.8.3) has been added. Perform the test starting with 9.3 and calculate the recovery. Ensure that the recovery is between 80 % and 110 %

9.3 	Extraction procedure

Cool the sample to about 10 °C, if necessary, to prevent losses of the extracting agent by volatilization.

Acidify the sample to pH 2 by adding mineral acid (6.5), if this has not been done in the field (see clause 8).

Add about 80 g of magnesium sulfate (6.4) per 900 ml of sample to avoid emulsions.

NOTE 1 The addition of magnesium sulfate is not necessary if samples are known not to form emulsions.

Add 50 ml of extraction solvent standard solution (6.11.2) and a magnetic stirrer bar, close the bottle and stir vigorously for 30 min on the magnetic stirrer (7.10).

Remove the stopper and replace it by the micro separator (7.8).

Add enough water (6.1) to allow withdrawal of the extracting-agent layer from the micro separator, transfer it to the clean-up column (7.9) and proceed according to 9.4.

When transferring the organic phase to the clean-up column, take care to avoid the transfer of water, as this will incrust the surface of the sodium sulfate. It is recommended to transfer the organic layer in several steps using a pipette, or, when using the micro separator (7.8) (Figure A.2), to position the meniscus below the cock.

In the case of strong emulsions, centrifuge the extract as follows. Transfer the extracting-agent phase together with the emulsion into a 100 ml centrifuge tube (7.7) and close the tube. Break the emulsion by centrifuging the extract.

NOTE 2 Centrifugation for 10 min to 15 min is usually sufficient.

Extraction by shaking may be adequate, but verify the efficiency.

9.4 	Clean-up procedure

Transfer the n-pentane phase (see 9.3) to a small column (see Figure A.1) filled with 2 g of Florisil (6.7) and covered with a layer of 2 g of sodium sulfate (6.3).

NOTE 1 Pre-rinsing of the column with a few milliliters of n-pentane (6.2) can be useful to prevent the formation of channels.

Let the n-pentane phase (see 9.3), followed by an additional 10 ml of n-pentane (6.2), run through the column into a 50 ml volumetric flask until the liquid reaches the 50 ml mark.

NOTE 2 Other procedures using the same amount of Florisil for cleaning the extract, e.g. shaking the total extract with 2 g of Florisil during 30 min on a shaking apparatus, may be used as an alternative, provided the results are equivalent to the column testing of Florisil  (see 9.6).

9.5 	Concentration

External apparatus for concentrating the purified n-pentane extract is not allowed. The gas chromatograph (7.2) is equipped with an injection system that allows an injection of up to 100 (l of n-pentane. Injection of 50 (l is only required for low mineral oil contents, i.e. at levels near the limit of determination.  Depending on the range of concentrations normally measured, a fixed injection volume, e.g. 25 or 50 (l, shall be used both for the calibration and for the measurements of the extracts.

Leave the empty sample bottle to drain for 5 min. Close the bottle with the cap used previously and determine its mass (m2) to an accuracy of 1 g.

9.6 	Suitability testing of Florisil

Check the suitability of Florisil at regular intervals and each time a new batch of dried Florisil is used, as follows:

For this purpose, use a test solution of stearyl stearate (6.12) and a calibration solution of mineral oils (6.8).

Perform the clean-up procedure (9.4) with 10 ml of the stearyl stearate solution, then add extracting agent (6.11.2) to a volume of 25 ml. Transfer an aliquot of the purified solution to a septum vial and analyze by gas chromatography (9.7). Measure the peak areas of stearyl stearate after Florisil treatment. Dilute 0,5 ml of stearyl stearate solution with extracting agent (6.11.2) to 25 ml and analyze by gas chromatography. Calculate the ratio of the peak areas for stearyl stearate in the treated and in the untreated solutions. This ratio should be less than 1. If not, activate the Florisil according to 6.7.

Perform the clean-up procedure (9.4) with 10 ml of 2 mg/ml calibration solution of mineral oils (6.8), then add extraction solvent standard solution (6.11.2) to a volume of 25 ml. Transfer an aliquot of the purified solution into a septum vial and analyze by gas chromatography.

Determine this mineral oil recovery on the basis of the peak area between C10 to C40 in the treated (with Florisil) and untreated calibration solutions. It should be at least 80 %. If this criterion is not met, wash out the Florisil batch with ample water and activate the Florisil as given in 6.7. If the repeated test shows again that the criterion is not met, try another batch of Florisil with a different batch number.

9.7 	Determination by gas chromatography

9.7.1 	Adjusting the gas chromatograph

Select a capillary column with one of the stationary phases specified (see 7.3 and annex B) for gas chromatographic analysis. Adjust the gas chromatograph until optimum separation is obtained. The peaks in the gas chromatogram of the standard mixture of n-alkanes (6.9) shall be baseline-separated. The relative response (peak area) of n-tetracontane (C40H82) compared with n-eicosane (C20H42) should be at least 0,8. If not, the discrimination of the injection system is too high and the injection system shall be optimized or replaced.

9.7.2 	Calibration

9.7.2.1	General

For calibration a distinction is made between initial calibration, working calibration and checking of the validity of the calibration curve. Initial calibration determines the working range and the linearity of the calibration function according to ISO 8466-1. Perform this calibration when the apparatus is used for the first time.

In the next step establish the final working range and perform the routine calibration. Repeat this calibration after maintenance (e.g. replacement of the capillary column), after repair of the gas chromatographic systems, and in case the system has not been in use for a longer period of time, or if the validity criteria cannot be met. Check the validity of the initial calibration with each series of samples to be analyzed.

Correct all chromatograms for column bleeding, for this purpose run blank chromatograms (chromatograms with solvent only) and use these for baseline correction.

9.7.2.2	Initial calibration

Establish the preliminary working range by analyzing at least five dilutions of the calibration standard mixture (6.8). Test for linearity in accordance with ISO 8466-1.

9.7.2.3	Routine calibration

After examining the final working range, analyze a minimum of five dilutions of the standard calibration mixture (6.8). Calculate a calibration function by linear regression analysis of the corrected peak areas. The actual sensitivity of the method may be estimated from the calculated regression function.

9.7.2.4	Check of the validity of the calibration function

Check the validity of the calibration function from the routine calibration with each batch of samples by analysis of one standard solution after every ten samples. The concentration of this standard solution shall lie between 40 % and 80 % of the working range. Make sure that the individual results do not deviate by more than 10 % of the working calibration line. If this criterion is met, assume the calibration to be valid. If not, recalibrate in accordance with 9.7.2.3.

For large batches of samples the number of analyses of the standard solution may be reduced, provided that at least three measurements are obtained for calculating the mean result.

9.7.3 	Measurement

Measure the sample, the calibration solutions and the blank solution in the gas chromatograph.

At regular intervals record a gas chromatogram (blank chromatogram) by injecting extracting agent (6.2) and analyze under the same conditions as for the sample. Use a chromatogram of extracting agent to correct the areas of the chromatograms of the samples.

NOTE 	An increase in "column bleeding" may indicate contamination of the injection system or the column.

9.7.4 	Integration parameters

Integrate the gas chromatogram between n-heptane and n-tetracontane. Start the integration just after the n-heptane peak at the signal level in front of the solvent peak (S in Figure B.3). End the integration just before the beginning of the n-tetracontane peak on the same signal level (E in Figure B.3).

Check all chromatograms visually to ensure correct integration. Draw a straight line from S to E. Mark the beginning and end of the integration on the chromatogram.

Identify the peaks of the aromatic hydrocarbons toluene, ethyl benzene and the three isomers of xylene. Integrate these peaks separately and subtract their peak areas from the total integrated area. For purposes of integration and subtraction these peaks should be assumed to extend to the baseline of the signal level

NOTE 1 For examples of chromatograms, see annex B.

The presence of peaks between solvent peak and n-heptane indicates that the sample probably contains low-boiling, volatile hydrocarbons.

Discrete peaks or an increased level of the baseline at the end of the chromatogram (retention time greater than that of n-tetracontane) indicate that the sample probably contains hydrocarbons with a high boiling point. This should be mentioned in the test report.

NOTE 2 The range of the carbon numbers of n-alkanes present in the sample is determined by comparing the gas chromatogram of the sample extract with that of n-alkane standard solution (6.9). The corresponding boiling range can be derived from annex C.

9.8 	Calculation

Calculate the hydrocarbon oil index using the equation:

(Am-b) . f . V . w

ρ =	-------------------

a . (m1 - m2)

where

ρ			is the content of dispersed mineral oil, in milligrams per liter;

a			is the slope of calibration function, in liters per milligram;

Am	is the integrated peak area of the sample extract, in instrument-dependent units;

f			is any dilution factor of the sample extract;

m1	is the mass of the filled sampling bottle, in grams;

m2	is the mass of the empty sample bottle, in grams;

w		is the density of the water sample, in grams per milliliter (for fresh water 1,00 g/ml may be used);

V		is the volume of the final extract, in milliliters;

b			is the intercept of the y-axis in instrument-dependent units.

9.9 		Expression of results

Express the concentration of mineral oil in water as hydrocarbon oil index, in milligrams per litre, to two significant figures.

EXAMPLES

Dispersed mineral oil	15 mg/l

Dispersed mineral oil	2,9 mg/l

# 10 	Test report

The test report shall refer to this document and contain the following details:

a) 	identity of the sample;

b) 	the content of dispersed mineral oil, in milligrams per litre;

c) 	any peculiarities observed during the test;

d) 	any actions not specified in this document, which may have influenced the result.

In addition, the following qualitative information can be provided in the test report:

e)	the boiling range of the mineral oil detected on the basis of the relative retention time versus the boiling point of the calibration mixture of n-alkanes;

f)	any presence of volatile hydrocarbons;

g)	any presence of hydrocarbons with a high boiling point.

# 11	Precision

To date no precision data is included for the OSPAR Reference Method, though a study has been carried out by NEL/LGC, involving a number of laboratories in the UK, the Netherlands, Norway and Denmark. The report for which can be found at <website link to be added>. The study indicated that the repeatability and reproducibility of the method were 15% and 20% respectively which is applicable up to a maximum dispersed oil in water concentration of 43mg/l.

In addition the instrumental part of this method has been checked in a small ring test.

Four laboratories in four countries (Denmark, the Netherlands, Norway and the United Kingdom) have analysed two standard solutions of hydrocarbons in the range C7 – C40 with different levels of toluene, ethyl benzene and xylenes. Each laboratory analysed the standards in triplicate on two different days.

The results are given in table 11.1, expressed as the ratio between the measured and the true values

Table 11.1: Ring test results for the OSPAR reference method

|  | Day | Lab 1 | Lab 2 | Lab 3 | Lab 4 |
| --- | --- | --- | --- | --- | --- |
| Low TEX | 1 | 0.95 | 1.01 | 0.98 | 1.21 |
| Low TEX | 2 | 0.93 | 1.01 | 0.87 | 1.09 |
| High TEX | 1 | 0.96 | 1.06 | 1.01 | 1.35 |
| High TEX | 2 | 0.92 | 1.02 | 0.90 | 1.13 |

In addition the laboratories were asked to determine also the content of mineral oil according to ISO 9377-2, i.e. only the hydrocarbons in the range C10 – C40. These results are given in table 11.2, expressed as the ratio between the measured and the true values

Table 11.2: Ring test results for ISO 9377-2

|  | Day | Lab 1 | Lab 2 | Lab 3 | Lab 4 |
| --- | --- | --- | --- | --- | --- |
| Low TEX | 1 | 0.97 | 1.04 | 1.06 | 1.31 |
| Low TEX | 2 | 0.93 | 1.03 | 0.93 | 1.18 |
| High TEX | 1 | 0.97 | 1.08 | 1.08 | 1.43 |
| High TEX | 2 | 0.91 | 1.04 | 0.96 | 1.19 |

The results of Lab 4 differ from the results of the other laboratories. Lab 4 did not use a large volume injection system, resulting in a too low signal.

The results of the other three laboratories show that modification of ISO 9377-2 by extending the range of hydrocarbons to C7 – C40 and excluding the aromatic hydrocarbons toluene, ethyl benzene and the isomers of xylene, have little or no effect compared to ISO 9377-2.

# Annex A – Example of a column

(informative)

Dimensions in millimeters

Key

1	Glass bulb

2	Fritted glass filter

Figure A.1 – Example of a column

# Annex B Examples of gas chromatograms of mineral oil standard and water samples

(informative)

The gas chromatograms have been recorded under the following conditions:

Injection technique:	programmed temperature vaporization (PTV)

Injection temperature:	50 °C to 300 °C

Injection volume:	1 μl

Column length:	30 m

Column internal diameter:	0,25 mm

Liquid phase:	DB 5 MS

Film thickness:	0,25 μm

Pre-column:	deactivated fused silica capillary

Carrier gas:	hydrogen

Carrier gas pressure:	0,8 bar

Oven temperature programme:	40 °C for 5 min, 10 °C/min to 300 °C, 300 °C for 20 min

Detector:	flame ionization detector

Detector temperature:	300 °C

Make-up gas:	nitrogen

Make-up gas flow:	25 ml

Key

1	n-Heptane

2	n-Tetracontane

Figure B.1 – Gas chromatogram of a standard solution (0,5 mg/l)

Key

1	n-Heptane

2	n-Tetracontane

Figure B.2 – Gas chromatogram of a surface water contaminated with fuel oil

Key

1	n-Heptane

2	n-Tetracontane

Figure B.3 – Gas chromatogram of a waste water

Key

1	n-Heptane

2	n-Tetracontane

Figure B.4 – Gas chromatogram of a waste water contaminated with lubricating oil

Example of the instrumental conditions for the determination of dispersed oil in produced water by gas chromatography

Apparatus and conditions

- Gas chromatograph:		Trace GC 2000

- Precolumn: 	10m * 0.53mm Fused silica coupled with 
5m * 0.53mm * 0.5μm CP-SIL 8CB

- Analytical column:		CP-SIL 5CB (25m * 0.32mm * 0.4μm)

- Carrier gas:			Helium

- Flow rate:			3.0 ml/min (constant flow)

- Injection volume:		50 μl

- Injection speed:			4 μl/sec

- SVE time:			0.05 min

- Extraction liquid: 		n-pentane

Temperature programme

n-Heptane and benzene (if possible) shall be fully separated from the extraction solvent n-pentane.

The following programme has been used:

Fig B.5 Temperature programme

Identification

Solutions of individual BTEX-compounds have been used for allocation of retention times.

Fig. B.6  Standard solution of  BTEX

Fig B.7 Calibration mixture of Mineral Oil, including C7, C10 and C40 internal standards

(Calibration mixture from National Institute for Public Health and the Environment – The Netherlands)

Fig B.8 	Overlay of standard solution of BTEX and calibration mixture of Mineral Oil including C7, C10 and C40 internal standards

# Annex C Determination of boiling range of a mineral oil from the gas chromatogram

(informative)

Table C.1 shows the boiling points of the n-alkanes containing between 2 and 44 carbon atoms (up to and including C44). Using these boiling points, it is possible to determine the boiling range of the mineral oil by comparing the gas chromatogram of the calibration mixture of n-alkanes with that of the sample extract.

Table C.1 - Boiling points of n-alkanes

| Number of carbon atoms | Boiling point °C | Number of carbon atoms | Boiling point °C |
| --- | --- | --- | --- |
| 2 3 4 5 6 7 8 9 10 11 12 13 14 15 16 17 18 19 20 21 22 | -89 -42 0 36 69 98 126 151 174 196 216 235 253 271 287 302 317 331 344 356 369 | 23 24 25 26 27 28 29 30 31 32 33 34 35 36 37 38 39 40 41 42 43 44 | 380 391 402 412 422 432 441 450 459 468 476 483 491 498 505 512 518 525 531 537 543 548 |

# Bibliography

[1]	EN 590:1998, Automotive fuels - Diesel - Requirements and methods of test.

[2]	KVERNHEIM A.L. et al., Development of a new hydrocarbon index for oil-in-water. Chemosphere, 39, 1999, pp. 2707-2722.

# Footnotes

Amendments to this Agreement were adopted by OIC 2011. See OIC 11/13/1, paragraph 2.10

Standards are available from

| Fachgruppe 1.2 Bundesanstalt fur Materialforschung and -prüfung Richard-Willstätter-Strasse 11 D-12489 Berlin, Germany | Chiron AS Stiklestadveien 1 NO-7041 Trondheim Norway | Sigma-Aldrich Company Ltd. The Old Brickyard New Road Gillingham Dorset SP8 4XT UK |
| --- | --- | --- |

This information is given for the convenience of users of this document and does not constitute an endorsement by OSPAR of these products.